Specification for laptop battery tester RFNT3

Summary	2

Basic Parameters	2

Function Introduction	4

Charge	4

Discharge	4

Test	5

Capacity Correction	5

50% Capacity	6

Online Functions	6

Monitoring	6

Control	6

Parameters	6

Attention	6

Others	7

Structure	8

Accessories	8

Front Panel	9

Back Panel	10

Top Panel	10

Bottom Panel	10

Display Mode	11

LCD Display	11

LED Indication	12

Operation	12

Power Up	12

Connecting Battery	12

Charge	12

Discharge	13

Function Menu	13

Online Functions	13

Cautions	14

# Summary

RFNT3 is a multi-function battery testing device, having the features below:

- with one channel, supporting one intelligent battery

- Supporting charge, discharge, test, capacity correction, charge / discharge 50% capacity

- LCD for displaying battery information

- LED for indicating working status

- With charge-key, discharge-key, backlight-key and function-key

- Pre-charge, fast charge and constant voltage charge for charging stages and the variation range of the values of the constant current and the constant voltage is less than 1%

- Constant current for discharging

- With a USB port for connecting with computers, master machine can monitoring and control slave machine

# Basic Parameters

Dimensions: 201 x 99 x 35 mm

Weight: 333g

Adapter    Input: AC 100 ~ 240 V   Output: DC 24V 2.65A

Display mode

Segment LCD with white backlight + LED in red and green

Discharging mode

Constant current discharge for double loads, inner load & external load

Charging mode

Constant current and constant voltages are the main mode. Active charge and pre-charge are available according to different battery voltage.

Active charge:

Tester may not recognize 0V-battery. Users should press the charge key to start up the active charge function. Being activated and going to normal charging status, when the battery establishes communication with the tester.

Pre-charge:

Tester goes into pre-charge status when voltage of single cell < 3000mV.

Constant current:

Tester goes into fast charge status when voltage of single cell > 3000mA & < 4200mA. Fast charge current relates to the battery nominal capacity.

At the end of charge:

The current intensity begins to decrease when tester goes into constant voltage charge status. It stops charging when the capacity reaches 100% and the current intensity less than 1A or the current intensity reduce to the end charge current.

Charging current

There is a connection between current intensity and charge mode. The details as below:

The max charge current is 2.5A.

Nominal capacity Pre-charge current Fast charge current End charge current

≥6000                400mA            2500mA              300mA

≥3500                400mA            2000mA              200mA

Else                  200mA            1000mA              100mA

Standby power consumption≤1W

Working temperature

0~40℃               20%－90% RH (Non condensation)

Storage temperature

-20℃~65℃           10%－95% RH (Non condensation)

Time

Time limit for charging:                  8.8h

Time limit for discharging:                8.8h

Time limit for testing:                    6s + 6s = chg + disc

Auto quit time for function menu:          10s

# Function Introduction

## Charge

Active, pre-charge, constant current and constant voltages are the stages of charging mode. The initial charging mode depends on the battery voltage, so active and pre-charge don't happen to every batteries. For charging mode and current, please refer to the paragraphs of charging mode and charging current on page 2 & 3.

Tester will stop charging if one of the conditions below was met.

- Remaining capacity  ≥ 99% & current intensity = 0

- Remaining capacity = full charge capacity

- Current intensity reduce to end charge current when constant voltage charge

- Remaining capacity reach up the set value

## Discharge

Constant current discharge for double loads

Inner load: 0 ~ 1000mA, adjustable continuously via Battery Manager

External load: 0mA, 500mA, 1000mA, 1500mA, 2000mA, five gears with dip switches offered

Tester will stop discharging if one of the conditions below is met.

- Remaining capacity ≤set value

- Voltage ≤set value

- Current intensity reduce to zero

Remarks:

Many batteries can not measure exact current intensity, which cause an else current value appears after you set a value. However, the actual discharging current is the same with your set value.

## Test

Testing time is 12 seconds and the purpose is to check both the charging and discharging performances. After finishing test, LED flickers with green color and LCD reports the test result.

Normal:              Test           OK            Shown on LCD

Abnormal:            Test           NG            Shown on LCD

Possible causes and solutions for the abnormal:

- Cause: battery voltage is too low to discharge

Solution: try to charge the battery to raise voltage

- Cause: remaining capacity is too full to charge any more

Solution: try to discharge the battery to reduce the remaining capacity

- Bad contact of battery connecting wires

- Damaged battery

- Tester trouble

## Capacity Correction

The battery capacity value in the battery management chip will different from the actual capacity after a battery goes through many recycling cycles. It will affect the use of the battery when the difference reaches a certain degree. Capacity correction is to deep discharge a battery and then fully charge, which will correct the capacity value in the battery management chip.

Remarks:

Discharging current can be set via Battery Manager. Deep discharge will cause a battery losing communication with tester, so battery dropping phenomenon may happen in the discharging process if the end capacity and voltage are set too low. In such a case, active charge will make the battery work again.

## 50% Capacity

CD50%: charge firstly, and then discharge to 50% capacity

DC50%: discharge firstly, and then charge to 50% capacity

## Online Functions

Tester can be connected with master machine via a USB port. Users can monitor, control and set parameters via Battery Manager software in the master machine.

## Monitoring

Battery information can be monitored via three windows.

- Home screen

Battery bar and brief information can be read on home screen.

- Battery detail information window

Show smart battery data in real time and the data can be saved and printed in image format.

3.  Battery graph window

Voltage, current and capacity are traced in curves.

## Control

Battery Manager can control slave machine completely. Right-click menu on the battery bar includes the related commands.

## Parameters

Users can set working parameters via set parameter window in Battery Manager.

## Attention

All the functions of charge, discharge, capacity correction and 50% capacity are comply with the charging discharging conditions, which are showed on function introduction paragraph in page 4 & 5, so parameters should be set reasonably. Especially, deep discharge will cause a battery losing communication with tester.

## Others

- Active cooling

- Active batteries of zero voltage

When a battery is deep discharged, the discharge protection circuit will start up and the voltage in the battery connector will become zero voltage. In this case, an appropriate voltage should be supplied to active the battery.

- Battery identification

Battery information can be read from battery management chip after establishing communication.

- Intelligent protection

Against over-current, over-voltage, over-charge, over-heat and short circuit.

- Fan

Discharging mode: fan startup

Other modes: fan startup when temperature above 48℃

Fan stop working when temperature below 40℃

- Trouble alarm

LED flickers in red color and current command stops when one of the conditions below happen.

1). There is not current for around 50 seconds when charge.

2). There is not current for around 25 seconds when discharge.

3). Charging time is above 8.8 hours.

4). Discharging time is above 9.9 hours.

5). Suitable for almost all laptop batteries, above 95%

# Structure

## Accessories

1).1 * adapter

2). 1 * USB cord

3). 1 * CD

4). 13 * connectors

5). 10 * cables

6). 10 * jumpers

7). 1 * load

## Front Panel

- Segment LCD

- LED

- LED

- KeyFun

- KeyDisc

- KeyChg

- KeyBL & KeyOK

KeyBL: used for controlling LCD backlight

KeyOK: used for choosing a command in function menu

## Back Panel

- Heat dissipation holes

## Top Panel

- AC input jack

- Heat dissipation holes

- mini USB port

## Bottom Panel

- Connecting wire jack for connecting to a battery

- Heat dissipation holes

- Connecting wire jack for connecting with external load

# Display Mode

## LCD Display

- Battery voltage

- Remaining capacity

- Charging marks

- Charging current or discharging current

- Discharging marks

- Remaining capacity percentage

- Charge / discharge cycle count

- Full charge capacity or Function number

- Test marks

- Test successful

Test failure

Battery bar

## LED Indication

Shine in red color:     On discharge

Flicker in red color:    Trouble alarm

Shine in green color:   On charge

Flicker in green color:  A task done

# Operation

## Power Up

- Connect the adapter to AC power source

- Connect the tester with the adapter

LCD shows battery information and LED lights up for a moment if a battery has been connected to the tester.

Both LCD backlight and LED lights up for a moment if no battery has been connected to the tester.

## Connecting Battery

1). Make sure the connecting wire fits the battery

2). Make sure the connecting wire is plugged into the battery socket in correct direction

3). Make sure no poor connections

## Charge

- Normal charge

Tester starts to charge your battery after you press the charge button. Battery information will be showed on LCD screen in real time.

- Active charge

If the battery needs to be activated, tester can not show any battery information. In this case, you should press the charge button to start up active charge. Tester will go into normal charge process after the battery is activated successfully.

Time limit of active charge:             40 seconds

Output voltage of active charge:          8000mV

Output current of active charge:          200mA

## Discharge

Tester starts to discharge your battery after you press the discharge button. Battery information will be showed on LED screen in real time.

## Function Menu

Keyfun button comes with five menus and the details as below.

F1   Test

F2   Capacity correction

F3   Stop

F4   CD50%

F5   DC50%

Shift menus:                 press KeyFun button

Choose command:            press KeyOK button

When the related command is being done, the function number will show on LCD screen and you can have a look at other function numbers by pressing KeyFun button.

For detailed functions of above commands, please refer to the paragraphs of function introduction on Page 4 & 5&6

## Online Functions

- Connect tester to your computer via USB cord

- Startup Battery Manager

- Refer to Battery Manager User Manual

Remarks: Battery Manager software and USB drive program should be installed in your computer before you use this function.

# Cautions

- Use the fitted adapter. Nonstandard adapters might lead to working disability.

- To avoid explosion and danger, damaged batteries and leaking batteries are not allowed.

- The product is good for indoors. Do not expose and use it in outdoors.

- Do not drop or put something heavy on it.

- Children are not allowed to use the product without adults.

- Do not keep the product under the direct rays, beside the hot devices or in other hot environment.

- Do not disassemble or assemble the product.